- Measurement of Repeatability and Uncertainty:

| Sr. No. | Standard weight | Calibration Start Time | Acceptance criteria | Observed weight |  |  |  |  | Calibration End Time | Uncertainty = (SD x 2) / Std. Wt. (NMT 0.10 %) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 150 kg |  | 149.85 to 150.15 kg | 1 | 2 | 3 | 4 | 5 |  |  |
|  |  |  |  | 6 | 7 | 8 | 9 | 10 |  |  |

Remarks: Satisfactory/Not Satisfactory

| Calibrated by: (Sign. /Date) | Checked by: (Sign. / Date) |
| --- | --- |

- Drift measurement:

| Sr. No. | Standard weight | Calibration Start Time | Acceptance criteria | Observed weight |  |  |  |  | Calibration End Time | Acceptance criteria |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 100 kg |  | 99.90 to 100.10 kg | 0 | 1 | 2 | 3 | 4 |  | Tolerance limit is ± 0.10% of actual weight. |

Remarks: Satisfactory/Not Satisfactory

| Calibrated by: (Sign. /Date) | Checked by: (Sign. / Date) |
| --- | --- |

- Whole Range Calibration:

| Sr. No. | Standard weight | Calibration Start Time | Observed weight | Acceptance criteria Tolerance limit is ± 0.10% of actual weight. | Calibration End Time | Status Ok / Not Ok | Remark |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 1.00 kg |  |  | 1.00 to 1.00 kg |  |  |  |
| 2 | 150.00 kg |  |  | 149.85 to 150.15 kg |  |  |  |
| 3 | 200.00 kg |  |  | 199.80 to 200.20 kg |  |  |  |
| 4 | 300.00 kg |  |  | 299.70 to 300.30 kg |  |  |  |

Remarks: Satisfactory/Not Satisfactory

| Calibrated by: (Sign. /Date) | Checked by: (Sign. / Date) |
| --- | --- |